## Slide 1: Fast timing of double sided Si strip detector
Lecture: Hans-Jürgen Wollersheim
e-mail: h.j.wollersheim@gsi.de
Akhil Jhingan, Inter University Accelerator Centre, Delhi

## Slide 2: Fast timing of double sided Si strip detector
Ideal beam tracking detector:
large area tracking system
Energy resolution ΔE/E < 1%
Position resolution ~ 1 mm
Time resolution ≤ 100 ps
RIB-experiments after slowing down to Coulomb barrier
double-sided silicon-strip detectors
- active area 50x50 mm2
- thickness 40 μm
- 16 x 16 strips (3.125 mm)
- manufactured by MICRON
MCP DSSSD

## Slide 3: Micro Channel Plates - MCP
time resolution – 150 ps
position resolution
for alpha particles – 3 mm
for fission fragments – 1.7 mm
N. Kondratiev (FLNR JINRDubna), M. Pfeiffer (IKP Köln)

## Slide 4: Micro Channel Plates - MCP
time resolution – 150 ps
position resolution
for alpha particles – 3 mm
for fission fragments – 1.7 mm
N. Kondratiev (FLNR JINRDubna), M. Pfeiffer (IKP Köln)

## Slide 5: Micro Channel Plates - MCP

## Slide 6: Schematic diagram of the MCP-based position-sensitive detector

## Slide 7: Tagging of triple alpha source

## Slide 8: Fast timing of double sided Si strip detector
Experimental ToF setup with MCP and DSSSD

## Slide 9: Fast timing of double sided Si strip detector
front strips
timing preamplifiers

## Slide 10: Fast timing of double sided Si strip detector
 | 229Th α-source | Cf fission source
rise time | 5 ns | 9 ns
pulse hight | 15 mV | 300 mV
noise | 5 mV | 
Energy resolution ~300 keV for different strips

## Slide 11: Fast timing of double sided Si strip detector
Energy measurement with Mesytec:
Energy resolution ΔE = 300 keV
229Th
Time measurement with 16-ch. fast timing preamplifier (IUAC):
pulse height 15 mV
noise 5 mV
rise time 5 ns
energy gated TOF spectra:
Eα=6.34 MeV
Eα=7.07 MeV
Δt ≈ 1 ns
MCP DSSSD

## Slide 12: Fast timing of double sided Si strip detector
ToF versus alpha energy measured for four strips
Details of the 229Th alpha source:
Eα = 4.84 MeV, v = 1.527 107 m/s, TOF = 9.824 ns
Eα = 5.83 MeV, v = 1.676 107 m/s, TOF = 8.950 ns
Eα = 6.34 MeV, v = 1.747 107 m/s, TOF = 8.586 ns
Eα = 7.07 MeV, v = 1.845 107 m/s, TOF = 8.130 ns
Eα = 8.38 MeV, v = 2.009 107 m/s, TOF = 7.466 ns
Energy gated ToF spectra

## Slide 13: DSSSD – fast timing
Estimated time resolution:
ΔT (FWHM) ~ 100 ps
Coulomb scattering of 48Ca beam, 12.6 MeV/u at 200

## Slide 14: Fast timing of double sided Si strip detector
Energy measurement with Mesytec:
Energy resolution ΔE/E = 1.5%
MWPC + MCP MCP + DSSSD
Time measurement with 16-ch. fast timing preamplifier (IUAC):
pulse height 300 mV
rise time 6.3 ns
TOF between MCP and DSSSD
Δt ≈ 1 ns
IUAC experiment :
28Si → 197Au (0.2mg/cm2) at 122 MeV
Time resolution 200 ps for one of the 256 detector pixels
one detector pixel
total detector